www.msti.co.kr

Clean Bench Controller

MEXION

Makes eXcellent Incense Of Nature

MECHATRONICS SYSTEM TECHNOLOGY

MST Co., Ltd.

◆ Table of Contents ◆

□ General Specifications for Korean Version

1-1  Temperature......................................................................................................................................1

1-2  Humidity............................................................................................................................................1

1-3  Output................................................................................................................................................1

□ Special Specifications for Korean Version

2-1  Protection of Setting .......................................................................................................................1

2-2  Fan Speed Control ...........................................................................................................................1

2-3  Filter Warning Time ........................................................................................................................1

2-4  Load Indicator ..................................................................................................................................1

□  Monitor of Korean version

...................................................................................................................2

□  Operation Instruction

4-1  Operation / Stop.................................................................................................................................3

4-2  Warning Information .........................................................................................................................4

4-3  Parameters ...................................................................................................................................5~6

□  Wiring Diagram

.........................................................................................................................7

□  Product List

.........................................................................................................................10

□  General Specifications for MICOM

-  -

Input Voltage	AC80V～230V FREE VOLTAGE
Output Voltage	AC 80V～230V
Terminal Capacity	SSR Output 3A
Internal Voltage	1500V 3 Min.

1-1  Temperature

‧ Temperature Scope for Operation :  5℃ ～ 65℃

1-2  Humidty

‧ Humidity Scope for Operation : 20 ～ 85%

1-3  Output Specifications

‧ FAN SPEED 3 POINT (LOW /MID /HIGH)

‧ F/ L  LAMP(Fluorescent light) 1 POINT

‧ U/V  LAMP(Germ-Killing Light) 1 POINT

□ Special Specifications for MICOM of Korean version

2-1  Protection of Setting

Setting values remain same for power failure.

2-2  Fan Speed Control

Easy to control fan speed during operation

2-3 Setting of Filter Warning Time

Change filter after full operation by Setting filter renewal time for filter durability life

2-4 Load Indication

Indication load of fan motor, fluorescent lamp. germ-killing lamp

□ Monitor of Korean version

[ Picture 1 ]  Korean Monitor (Vertical & Horizontal types)

● Power : Power button to start or stop

● U/V : UV button to operate germ-killing lamp

● ▲  : UP button to increase values for fan speed and time

● ▼  : DOWN button to decrease values for fan speed and time

(You can change parameters after 2-second warning sound by pressing on button)

● Fluorescent Lamp : Fluorescent lamp button to operate fluorescent lamp

(During setting parameters this button is used to change mode)

□ Operation Method

4-1 Start & Stop

●  Start

CBM-2000

Controller

Company Telephone

Company Name

[Picture 2] Initial Screen

When power is on the display shows machine name, customer name, and telephone number for 2 secondss like picture 2 .

Operation Time

0000 Hours

Company Telephone

Company Name

[Picture 3] Idle Time

During machine stands idle for stop status, the display shows company name, telephone number, operation time. Every outputs stop during idle time and the idle status remains in case of recovery from power failure.

Press the ON/OFF button to start at picture 3.

Operation Time

0000 Hours

Company Name

[Picture 4] Operation

You can control fan speed and  operation time is accumulated. You can select fluorescent lamp or erm-killing lamp.

● Stop

Operation Time

0000 Hour

Company Telephone

Company Name

[Picture 5] Stop of Operation

If you want to stop every output during operation,  press operation/stop button then the display shows   picture 5.

You can operate fluorescent lamp or germ-killing lamp by pressing buttons while idle time.

4-2  Warning Notice

Operation Time

0000 Hours

Company Telephone

Company Name

[Picture 6] Display for normal operation

Above display is for normal operation and the display shows warning information up to two in case of  warning.

Warning Notice

1.FAN Warning

2.Filter Warning

[Picture 7] Warning Notice

1. Fan warning with alarm when engine doesn't have load while operation of fan.

2. Filter warning with alarm when filter life is up.

* You can stop the alarm by pressing UP or DOWN button during warning.

4-3 Setting Parameters

Alarm sounds for 2 seconds when you press operation button in idle.

The you can change the display to parameter setting mode by pressing DOWN button while alarm sounds.

1)CT Sensor  1: 000

2)CT Sensor  2: 000

3)CT Sensor 3: 000

4)Recovery from Power Failure : 000 Sec.

[Picture 8] FAN Load (A)

[Picture 8] The display shows fan load. You can go to next menu by pressing fluorescent lamp button.

1)CT Sensor1: 000

2)CT Sensor  2: 000

3)CT Sensor 3: 000

4) Recovery from Power Failure: 000 Sec.

Picture 9] U/V(Germ-killing) Load (A)

[Picture 9] The display shows germ-killing lamp load.

You can go to next menu by pressing fluorescent lamp button.

1)CT Sensor 1: 000

2)CT Sensor2: 000

3)CT Sensor3: 000

4) Recovery from Power Failure : 000 Sec.

[Picture 10] The display shows fluorescent lamp load.

You can go to next menu by pressing fluorescent lamp button.

1)CT Sensor1: 000

2)CTSensor2: 000

3)CT Sensor3: 000

4)Recovery from power failure: 000 Sec.

[Picture 11] Recovery from Power Failure

[Picture 11] Menu is to set recovery time from power failure. (Setting value: 0 to 240 sec.)

You can go to next menu by pressing fluorescent lamp button.

5)Set Time: 0 hrs.

6)Operation Time: 0 hrs.

7)UV     :  Stop

8)Fluorescent lamp : Stop

[Picture 12] Setting operation time

[Picture 12] Menu is to set operation time according to filter durability life. The display shows filter change time when total operation time reaches filter durability time. You can go to next menu by pressing fluorescent lamp button.

5)Set Time: 0000 hrs.

6)Operation Time: 0 hrs

7)UV     :  Stop

8)Fluorescent lamp : Stop

[Picture 13] Operation Time

[Picture 13] The display shows total operation time.

You can go to next menu by pressing fluorescent lamp button.

5)Set Time: 0000 hrs.

6)Operation Time: 00 hrs.

7)UV     :  Stop

8)Fluorescent lamp:  Stop

5)Set Time: 0000 hrs.

6)Operation Time: 00 hrs.

7)UV     :  On

8)Fluorescent lamp:  Stop

[Picture 14] UV (Germ-killing) lamp operation check

[Picture 14] Menu is to operate germ-killing lamp. The germ-killing lamp is on when you press "UP" button one time in idle status. You can go back to idle status by pressing "DOWN" button. You can go to next menu by pressing fluorescent lamp button.

5)Set Time: 0000 hrs.

6)Operation Time: 00 hrs.

7)UV     :  Off

8)Fluorescent lamp: Stop

5)Set Time: 0000 hrs.

6)Operation Time: 00 hrs.

7)UV     :  Off

8)Fluorescent lamp: On

[Picture 15] Fluorescent Lamp check

[Picture 15] Menu is to operate fluorescent lamp. The fluorescent lamp is on when you press "UP" button one time in idle status. You can go back to idle status by pressing "DOWN" button. You can go to next menu by pressing fluorescent lamp button.

9)FAN Warning: On

9)FAN Warning: On

[Picture 16] FAN Warning check

[Picture 16] Menu is to set fan warning. You can change "on" to "off" or vice versa by pressing "UP" button or "DOWN" button. The controller doesn't check fan status if you set it with "off".

* Setting values are automatically stored by pressing power button when setting is completed.

□  Wiring Diagram

8	7	6	5
4	3	2	1

1.FAN SPEED COMMON (Brown)

2.U/V (Violet)

3.Fluorescent (Gray)

4.Earth (Green)

5.FAN SPEED COMMON (White)

6.FAN SPEED LOW (Red)

7.FAN SPEED MID (Blue)

8.FAN SPEED HIGH (Black)

□ Products List

1. Controller for Temperature & humidity (Thermo-hygrostat)

2. Temperature controller (Air conditioner)

3. F‧F‧U CONTROLLER

4. Dry Cooler  CONTROLLER

5. Cpu Chiller CONTROLLER

6. Controller for mushroom (Dry - Wet type)

7. Controller for low temperature storage

8. Thermo-hygrostat controller

9. Hygrostat controller

10. Controller for refrigerator (Standard, premium type)

11. Remote control system for thermo-hygrostat

12. Remote controller for mushroom

13. Remote control for FFU  (Group or individual control)

※ The machine can control temperature, humidity, pressure, carbon dioxide, and oxygen.

We can supply additional control items on demand.

✔ The product can be modified for improvement at any time.

✔ For replacement or repair of the product due to malfunction or breakdown, please contact the head company directly.

MST Corporation

198-6, Angok-ro, Bucheon-si, Gyeonggi-do, Republic of Korea

Tel. 032-343-2577~8 / Fax. 032-343-2579 / Mail. mst2577@hanmail.net / www.msti.co.kr